BANGLADESH UNIVERSITY OF ENGINEERING AND TECHNOLOGY

Course No: ChE 206

Course Title: Chemical Engineering Laboratory I

A Long Report on
Experiment No 04: CALIBRATION OF ORIFICE METER

Submitted To:
Mahmud Hayet Rudro
Lecturer, Department of Chemical Engineering, BUET

Submitted By:
Mohammad Shahriar Hassan Sifat
Student ID: 2302004
Section: A1
Lab Group: 01
Term/Semester: Level 2 Term Ⅱ

Lab Partners:
Ahmed Arham Auritra, Student ID: 2302001
Abdullah Al Amin, Student ID: 2302002
Sameeha Tabassum Semeen, Student ID: 2302003
Md. Nazim Ahmed Sayem, Student ID: 2302005

Date of Performance: August 1,2026
Date of Submission: August 22,2026

Statement of Academic Integrity

In signing this statement, I hereby certify that the work on the “Calibration of Orifice Meter” is my own and that I have not copied the work of any other student (past or present) while completing this experiment. I understand that if I fail to honour this agreement and commit plagiarism, I will be disqualified and subject to possible disciplinary action.

Name: Mohammad Shahriar Hassan Sifat

Student ID: 2302004

Date: August 22, 2026

Signature

# Summary

In this experiment, an orifice meter was calibrated by studying the relationship between volumetric flow rate and the pressure-head difference developed across a sharp-edged orifice. Water was circulated through the system, and a differential U-tube manometer containing carbon tetrachloride was used to measure the pressure difference across the orifice. The actual flow rate was determined from the time required to collect a known volume of water, and eight observations were taken at different flow conditions. As the volumetric flow rate increased from 3.4941 × 10⁻⁴ to 7.4516 × 10⁻⁴ m³/s, the pressure-head difference increased from 0.041 to 0.153 m H₂O. The Reynolds number increased from 10586 to 22576, confirming that all observations were in the turbulent-flow region. The coefficient of orifice varied between 0.751 and 0.874, with an average value of about 0.804, showing reasonably consistent meter performance over the tested range. The pressure-head difference increased nonlinearly with flow rate, while the plot of ΔH  against flow rate was approximately linear with R² ≈ 0.962, supporting the theoretical relation Q ∝ ΔH. Although some variation in the discharge coefficient occurred because of pump fluctuations, manometer-reading uncertainty, air bubbles, timing errors, and possible slight rounding or wear of the orifice edge, the experiment successfully established a practical calibration relationship. Therefore, the calibrated orifice meter can be used to estimate unknown flow rates from measured pressure differences under similar operating conditions.

# Acknowledgement

I would like to express my sincere gratitude to the Department of Chemical Engineering, Bangladesh University of Engineering and Technology (BUET) for providing the necessary laboratory facilities and academic support to conduct the experiment entitled “Calibration of Orifice Meter” as a part of the course ChE 206: Chemical Engineering Laboratory I.

I am especially thankful to my experiment supervisor Mahmud Hayet Rudro, Lecturer, Department of Chemical Engineering, BUET, for his valuable guidance, clear instructions, and continuous supervision throughout the experiment. His explanations of the theoretical concepts and experimental procedures greatly helped in understanding the working principle and calibration of the orifice meter. I would also like to thank the laboratory instructors and staffs for their assistance in operating the equipment safely and for providing the necessary technical support during the experiment.

Finally, I am grateful to my group members for their cooperation, teamwork, and active participation during the experimental work, data collection, and analysis. Their support contributed significantly to the successful completion of this experiment.

Table of Contents

Summary	i

Acknowledgement	ii

1.Introduction	1

2.Theory	2

2.1 Orifice Meter and Working Principle	2

2.2 Types of Orifice Meter and Orifice Plates	3

2.3 Vena Contracta	4

2.4 Continuity Equation	4

2.5 Bernoulli's Equation	5

2.6 Differential Manometer Relation	6

2.7 Theoretical Discharge Through the Orifice	6

2.8 Coefficient of Orifice	7

2.9 Reynolds Number	8

2.10 Calibration of the Orifice Meter	9

3.Methods and Materials	10

3.1 Apparatus	10

3.2 Experimental Setup	10

3.3 Experimental Procedure	11

4.Observed Data	12

5.Calculated Data:	13

6.Result and Discussion	15

7.Conclusion	20

Reference	21

Nomenclature	22

Appendix	23

List of Illustrations

List of Tables

Table 1: Data table for calibration results of the orifice meter.	15

Table 2: Data table for observed volume, required time and manometer readings	12

Table 3: Table for calculation of volumetric flow rate, cross-sectional area, average velocity and Reynolds number	13

Table 4: Data table for the calculated values of pressure head, square root of pressure head, theoretical flow rate and discharge coefficient.	14

List of Figures

Figure 1: Schematic Diagram of an Orifice Meter	2

Figure 2: Different types of orifice meters and orifice plates based on construction and geometry[3]	3

Figure 3: Schematic diagram of orifice meter apparatus	10

Figure 4: Plot of Pressure Head Difference vs. Volumetric Flow Rate	16

Figure 5: Plot of square root of pressure head difference vs volumetric flow rate	16

Figure 6: Plot of discharge coefficient vs pressure head difference	17

Figure 7: Plot of discharge coefficient vs square root of pressure head difference	18

Figure 8: Plot of discharge coefficient vs Reynolds number	18

# 1.Introduction

Accurate measurement of fluid flow rate is important in many engineering and industrial operations, including chemical processing, petroleum handling, water-supply systems, and power generation. Reliable flow measurements are required for process control, material-balance calculations, safe operation, and efficient use of equipment. Among the different flow-measuring devices, the orifice meter is widely used because of its simple construction, low cost, absence of moving parts, and ease of installation.

An orifice meter is a differential-pressure flow meter consisting of a thin plate with a sharp-edged circular opening installed perpendicular to the flow direction. As the fluid passes through the restricted opening, its flow area decreases and its velocity increases. According to Bernoulli's principle, this increase in velocity is accompanied by a decrease in static pressure. Therefore, a pressure difference is produced between the upstream and downstream sides of the orifice. By measuring this pressure difference, the volumetric flow rate through the pipe can be estimated.

The theoretical relationship between pressure difference and flow rate is obtained from the continuity equation and Bernoulli's equation. However, these equations assume ideal flow and do not fully account for viscosity, turbulence, flow separation, jet contraction at the vena contracta, and other mechanical-energy losses. Consequently, the actual discharge differs from the theoretical value. A correction factor known as the coefficient of orifice, Cₒ, is introduced to account for these effects. Calibration is therefore necessary to determine the actual performance of a particular orifice meter under practical operating conditions.

In this experiment, water is passed through a sharp-edged orifice meter and the pressure-head difference across the orifice is measured using a differential U-tube manometer containing carbon tetrachloride (CCl4). The actual volumetric flow rate is determined from a known volume of water collected over a measured time interval at several flow conditions. These measurements are used to calculate the theoretical flow rate, Reynolds number, and coefficient of orifice and to examine the relationship between flow rate and pressure-head difference. The main objective is to calibrate the orifice meter and establish a practical relationship that can later be used to determine an unknown flow rate directly from the measured pressure difference.

# 2.Theory

## 2.1 Orifice Meter and Working Principle

An orifice meter is a differential-pressure flow-measuring device used to determine the volumetric flow rate of a fluid through a pipe. It consists of a thin plate containing a sharp-edged circular opening, called the orifice, installed perpendicular to the direction of flow. When fluid approaches the orifice, the available flow area suddenly decreases. Consequently, the fluid velocity increases while its static pressure decreases. The pressure difference produced between the upstream and downstream sides of the orifice is therefore related to the rate of fluid flow through the pipe.

Figure 1: Schematic Diagram of an Orifice Meter

In the figure 1[1] below a schematic diagram of an orifice meter is shown.

The operation of an orifice meter is mainly based on the continuity equation and Bernoulli's equation. The continuity equation describes the increase in velocity caused by the reduction of flow area, while Bernoulli's equation[2] relates this increase in velocity to the corresponding decrease in pressure. Thus, part of the pressure energy of the fluid is converted into kinetic energy as it passes through the restriction.

## 2.2 Types of Orifice Meter and Orifice Plates

Orifice meters can be classified according to the construction of the meter body and the design of the orifice plate. Based on construction, the common types include single chamber, dual chamber, double block and bleed, and integral orifice plate meters. Single and dual chamber meters are commonly used for general flow measurement, while the double block and bleed design provides better isolation and improves safety during maintenance. In an integral orifice plate meter, the orifice plate and pressure-tapping arrangement are combined into a single compact unit, making it suitable for installations where space is limited.

Figure 2: Different types of orifice meters and orifice plates based on construction and geometry[3]

Based on the shape and position of the opening, different types of orifice plates are used for different flow conditions. An eccentric orifice plate has an opening located away from the center and is suitable for fluids containing suspended solids, as it helps prevent sediment buildup. A segmental orifice plate has a segment-shaped opening and is generally used for fluids containing gas bubbles or solid particles. A conical orifice plate provides a smoother entrance to the flow and can help reduce energy loss. The sharp-edged orifice plate is the most commonly used type for clean fluids, especially under turbulent-flow conditions. Multi-hole orifice plates contain several smaller openings and are mainly used to reduce noise and vibration at high flow rates.

In this laboratory experiment, a standard sharp-edged concentric orifice plate was used. This type was selected because clean water was used as the working fluid, with no significant suspended solids or gas bubbles. The sharp edge produced a clear flow contraction and measurable pressure difference across the plate, making it suitable for studying the relationship between discharge and pressure-head difference. Its simple geometry also allowed the experimental flow rate to be easily compared with the theoretical value, which was useful for evaluating the calibration and performance of the orifice meter.

## 2.3 Vena Contracta

When a fluid passes through a sharp-edged orifice, the streamlines cannot change direction abruptly at the edge of the plate because of the inertia of the fluid. Consequently, the stream continues to contract for a short distance after leaving the orifice. The location at which the flowing jet reaches its minimum cross-sectional area is known as the vena contracta. At this point, the fluid velocity is maximum and the static pressure is minimum.

After the vena contracta, the jet gradually expands again to occupy the full pipe cross-section. This expansion is accompanied by flow separation, turbulence, eddy formation, and irreversible energy losses. Therefore, the actual discharge through an orifice is lower than the discharge predicted by an ideal-flow analysis. The effect of these non-idealities is taken into account through the coefficient of orifice.

## 2.4 Continuity Equation

For steady incompressible flow, conservation of mass requires that the volumetric flow rate remain constant between any two sections of the pipe. Therefore,

Q = A1v1 = A2v2

where,

Q = volumetric flow rate

A1 = cross-sectional area of the pipe

A2 = cross-sectional area of the restricted section

v1 = average fluid velocity in the pipe

v2 = fluid velocity through the restricted section

Because the flow area at the orifice is smaller than the pipe area, the continuity equation requires the fluid velocity to increase as it passes through the restriction. If D is the internal diameter of the pipe and Dₒ is the orifice diameter, the corresponding areas are

A₁ =  πD24

Aₒ = πDₒ24

From continuity,

v₁ = AₒA1v₂ = DₒD2v₂

## 2.5 Bernoulli's Equation

For steady, incompressible and frictionless flow between an upstream section and the restricted section, Bernoulli's equation[2] is written as

P1ρg+v122g+z1=P2ρg+v222g+z2

where,

P1, P2= static pressures at the upstream and restricted sections

ρ = density of the flowing fluid

g = gravitational acceleration

v1, v2= fluid velocities at the two sections

z1, z2= elevations of the two sections

For the horizontal pipe used in the experiment, z₁ = z₂. Therefore,

P1- P2ρg =v22-v122g

The pressure-head difference across the orifice is defined as

ΔH =P1-P2ρg

Hence,

ΔH =v22-v122g

This relation shows that the increase in fluid velocity produced by the restriction is associated with a decrease in static pressure.

## 2.6 Differential Manometer Relation

The pressure difference across the orifice is measured using a differential U-tube manometer. In this experiment, water is the flowing fluid and carbon tetrachloride (CCl4) is used as the heavier manometric fluid. If R is the measured difference between the two manometer levels, the equivalent pressure-head difference expressed in terms of water is

ΔH = R(ρCCl4- ρwρw)

or equivalently,

ΔH = RρCCl4ρw- 1

where,

R = difference in manometer levels

ρCCl4 = density of carbon tetrachloride

ρw = density of water

The observed manometer difference must therefore be converted into an equivalent head of water before it is used in the orifice-meter discharge equation.

## 2.7 Theoretical Discharge Through the Orifice

From the continuity equation,

v₁ = DₒD2v2

and therefore,

v₁² = DₒD4v22

Substituting this relation into Bernoulli's equation gives

ΔH =v221 - DₒD42g

Hence, the theoretical velocity through the orifice is

v₂ = 2gΔH1 - DₒD4

Since the theoretical volumetric flow rate is Qtheo = Aov2,

Qtheo = Aₒ 2gΔH1 - DₒD4

This is the ideal orifice-meter equation obtained by combining the continuity and Bernoulli equations. The actual volumetric flow rate during the experiment is determined independently by measuring a known volume of water over a measured time interval:

Qexp =Vt

where,

V = measured volume of water

t = corresponding collection time

## 2.8 Coefficient of Orifice

The theoretical discharge equation assumes ideal flow with no mechanical energy loss. In an actual orifice meter, energy is lost because of viscosity, turbulence, flow separation, jet contraction, and subsequent expansion. Therefore, the experimentally measured flow rate normally differs from the ideal theoretical value.

A correction factor known as the coefficient of orifice, Cₒ, is introduced to account for these effects. It is defined as the ratio of the experimental discharge to the theoretical discharge:

Cₒ =QexpQtheo ADDIN ZOTERO_ITEM CSL_CITATION {"citationID":"XbEIeCAH","properties":{"unsorted":false,"formattedCitation":"[4]","plainCitation":"[4]","noteIndex":0},"citationItems":[{"id":14,"uris":["http://zotero.org/users/local/yU6TYrnp/items/FNNUJSCT"],"itemData":{"id":14,"type":"book","abstract":"\"Fluid mechanics is an exciting and fascinating subject with unlimited practical applications ranging from microscopic biological systems to automobiles, airplanes, and spacecraft propulsion. Fluid mechanics has also historically been one of the most challenging subjects for undergraduate students because proper analysis of fluid mechanics problems requires not only knowledge of the concepts but also physical intuition and experience. Our hope is that this book, through its careful explanations of concepts and its use of numerous practical examples, sketches, figures, and photographs, bridges the gap between knowledge and the proper application of that knowledge\"--","edition":"Fourth edition","ISBN":"978-1-259-69653-4","language":"eng","number-of-pages":"1024","publisher":"McGraw-Hill Education","publisher-place":"New York, NY","source":"K10plus ISBN","title":"Fluid mechanics: fundamentals and applications","title-short":"Fluid mechanics","author":[{"family":"Çengel","given":"Yunus A."},{"family":"Cimbala","given":"John M."}],"issued":{"date-parts":[["2017"]]}}}],"schema":"https://github.com/citation-style-language/schema/raw/master/csl-citation.json"} [4]

Therefore,

Qexp= CₒAₒ 2gΔH1 - DₒD4

and

Cₒ = QexpAₒ 2gΔH1 - DₒD4

Since Qexp =Vt, the coefficient can also be written as

Cₒ =VtAₒ 2gΔH1 - DₒD4

For an ideal fluid, the coefficient would approach unity. In real flow, however, irreversible losses and jet contraction cause the measured coefficient to be lower. For a sharp-edged orifice, the coefficient depends on the geometry of the plate, Reynolds number, pressure-tap arrangement, and physical condition of the orifice edge. Therefore, the actual flow rate is normally lower than the theoretical flow rate. Qexp < Qtheo and thus Co < 1.

## 2.9 Reynolds Number

The Reynolds number is a dimensionless parameter used to characterize the flow regime and represents the relative importance of inertial and viscous forces. It is defined as

Re = ρvDμ

where,

ρ = density of the flowing fluid

v = average fluid velocity in the pipe

D = internal pipe diameter

μ = dynamic viscosity of the fluid

The average velocity in the pipe is obtained from

v =QA

The flow is generally classified as laminar for Re < 2000, transitional for approximately 2000 < Re < 4000, and turbulent for Re > 4000. Orifice meters are commonly operated in the turbulent region. At sufficiently high Reynolds number, the discharge coefficient becomes only weakly dependent on Reynolds number and remains approximately constant over the operating range.

## 2.10 Calibration of the Orifice Meter

Calibration is necessary because the theoretical discharge equation is based on ideal-flow assumptions, whereas real fluid flow contains viscous, turbulent, separation, and contraction losses. During calibration, the experimentally measured volumetric flow rate is compared with the corresponding pressure-head difference across the orifice for several different flow conditions. This establishes the actual relationship between pressure drop and flow rate for the specific meter.

From the discharge equation,

Q = CₒAₒ 2gΔH1 - DₒD4

For a particular meter operating over a range in which Cₒ remains approximately constant, all terms other than ΔH are constant. Therefore,

Q ∝ ΔH  [5]

or equivalently,

ΔH ∝ Q²  [5]

Thus, a plot of pressure-head difference, ΔH, against volumetric flow rate, Q, is expected to show a nonlinear, approximately parabolic relationship. In contrast, a plot of Q against ΔH (or ΔH against Q) should be approximately linear. This linear relationship provides the practical calibration curve of the orifice meter. Once the meter is calibrated, an unknown flow rate can be determined from the measured pressure-head difference without performing a separate volumetric collection for every measurement.

# 3.Methods and Materials

## 3.1 Apparatus

- Orifice meter

- Differential U-tube manometer

- Water reservoir

- Centrifugal pump

- Flow-control valve

- Flowmeter

- Stopwatch

- Thermometer

- Meter scale

## 3.2 Experimental Setup

The experimental setup consisted of a closed-loop water-circulation system containing a water reservoir, centrifugal pump, control valves, horizontal test pipe, orifice meter, differential U-tube manometer, flowmeter, and discharge line. Water was drawn from the reservoir by the centrifugal pump and delivered through the pipeline. A bypass line was provided to return excess water to the reservoir and to help regulate the flow through the main test section.

Downstream tapping

Outlet Pipe

Flow meter

Pressure control valve

Orifice plate

Bypass valve

Pump

Supply valve

Upstream tapping

Scale

Manometer Fluid

Manometer

Reservoir

Entrance valve

Main valve

Figure 3: Schematic diagram of orifice meter apparatus

A sharp-edged orifice plate was installed in the horizontal section of the pipe. Pressure tappings located on the upstream and downstream sides of the orifice were connected to a U-tube differential manometer containing carbon tetrachloride (CCl4) as the manometric fluid. The difference between the liquid levels in the two limbs of the manometer represented the pressure difference produced across the orifice. A graduated scale beside the manometer was used to record these readings.

The flow rate through the system was adjusted using the control valve. After the water passed through the orifice section, it flowed through the downstream measuring section and flowmeter before leaving through the discharge line. This arrangement allowed the pressure-head difference across the orifice and the corresponding actual volumetric flow rate to be measured at different operating conditions for calibration of the orifice meter.

## 3.3 Experimental Procedure

Before starting the experiment, the setup was inspected to ensure that there were no leaks, loose connections, or air bubbles in the manometer lines, and the room and water temperatures as well as the pipe and orifice dimensions were recorded. The water reservoir was then filled, the bypass valve was kept open, and the centrifugal pump was started to circulate water through the system. The flow-control valve was adjusted to obtain a steady flow rate, and sufficient time was allowed for the manometer levels to stabilize. The readings from both limbs of the differential manometer were then recorded, while the time required for a known volume of water to pass through the flow-measuring section was measured using a stopwatch. The control valve was subsequently adjusted to obtain different flow rates, and the corresponding manometer readings and collection times were recorded after steady conditions were reached. This procedure was repeated for all eight observations, after which the pump was switched off and the collected data were used to calculate the volumetric flow rate, pressure-head difference, fluid velocity, Reynolds number, theoretical discharge, and coefficient of orifice.

# 4.Observed Data

Nominal pipe diameter = 2 inch

Schedule no. = 40

Diameter of the pipe, D = 2.067 inch = 0.0525 m [6]

Orifice diameter, Do = 1 inch = 0.0254 m

Room temperature = 30℃

Water temperature = 30℃

Table 1: Data table for observed volume, required time and manometer readings

| Observation No. | Volume, V (L) | Time, t (s) | Manometer Reading |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Left (cm) | Right (cm) | Difference (cm) |
| 1 | 5 | 14.31 | 32.6 | 25.6 | 7.0 |
| 2 |  | 10.04 | 35.0 | 23.4 | 11.6 |
| 3 |  | 9.81 | 35.6 | 22.2 | 13.4 |
| 4 |  | 9.18 | 37.2 | 20.8 | 16.4 |
| 5 |  | 8.65 | 38.1 | 20.0 | 18.1 |
| 6 |  | 7.25 | 39.1 | 19.0 | 20.1 |
| 7 |  | 7.15 | 40.8 | 17.5 | 23.3 |
| 8 |  | 6.71 | 42.2 | 15.8 | 26.4 |

# 5.Calculated Data:

Internal diameter of the pipe, D = 2.067 inch = 0.0525 m [6]

Density of water at 30℃, ρw= 995.65 kg/m3 [7]

Density of CCl4 at 30℃, ρCCl4= 1574 kg/m3  [7]

Viscosity of water at 30℃, μw= 0.797 mPa.s  [7]

Table 2: Table for calculation of volumetric flow rate, cross-sectional area, average velocity and Reynolds number

| Number of observation | Volume, V (L) | Time interval, t (s) | Volumetric Flow rate, Q×104 (m3/s) | Cross section area, A×104 (m2) | Average velocity, v (m/s) | Reynolds number, Re |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 5 | 14.31 | 3.4941 | 21.6475 | 0.1614 | 10586 |
| 2 |  | 10.04 | 4.9801 |  | 0.2301 | 15088 |
| 3 |  | 9.81 | 5.0968 |  | 0.2354 | 15442 |
| 4 |  | 9.18 | 5.4466 |  | 0.2516 | 16502 |
| 5 |  | 8.65 | 5.7803 |  | 0.2670 | 17513 |
| 6 |  | 7.25 | 6.8966 |  | 0.3186 | 20894 |
| 7 |  | 7.15 | 6.9930 |  | 0.3230 | 21187 |
| 8 |  | 6.71 | 7.4516 |  | 0.3442 | 22576 |

The volumetric flow rate, average fluid velocity, and Reynolds number calculated from the raw observations in Table 1 are summarized in Table 2. Further analysis using relevant necessary equations yielded the theoretical volumetric flow rate and discharge coefficient, which are listed in Table 3.

Table 3: Data table for the calculated values of pressure head, square root of pressure head, theoretical flow rate and discharge coefficient.

| Number of observations | Manometer reading difference, R (cm CCl4) | Pressure head, ∆H (m H2O) | ∆H (m H2O)0.5 | Theoretical volumetric flow rate, Qtheo×104 (m3/s) | Discharge Coefficient, Co |
| --- | --- | --- | --- | --- | --- |
| 1 | 7.0 | 0.041 | 0.202 | 4.655 | 0.751 |
| 2 | 11.6 | 0.067 | 0.260 | 5.992 | 0.831 |
| 3 | 13.4 | 0.078 | 0.279 | 6.441 | 0.791 |
| 4 | 16.4 | 0.095 | 0.309 | 7.125 | 0.764 |
| 5 | 18.1 | 0.105 | 0.324 | 7.485 | 0.772 |
| 6 | 20.1 | 0.117 | 0.342 | 7.888 | 0.874 |
| 7 | 23.3 | 0.135 | 0.368 | 8.493 | 0.823 |
| 8 | 26.4 | 0.153 | 0.392 | 9.040 | 0.824 |

# 6.Result and Discussion

Table 4 summarizes the principal calibration results obtained from the eight flow conditions. The experimental volumetric flow rate increased from 3.4941 × 10-4 m3/s to 7.4516 × 10-4 m3/s, while the corresponding pressure-head difference increased from 0.041 m H2O to 0.153 m H2O. The square root of the pressure-head difference increased from 0.202 to 0.392 (m H2O)0.5. The Reynolds number ranged from 10586 to 22576, which places all observations well inside the turbulent-flow region. The discharge coefficient varied between 0.751 and 0.874, with an average value of approximately 0.804.

Table 4: Data table for calibration results of the orifice meter.

| Number of observation | Flow rate, Q×104 (m3/s) | Pressure head difference, ∆H (m H2O) | ∆H (m H2O)0.5 | Reynolds number, Re | Discharge Coefficient, Co |
| --- | --- | --- | --- | --- | --- |
| 1 | 3.4941 | 0.041 | 0.202 | 10586 | 0.751 |
| 2 | 4.9801 | 0.067 | 0.260 | 15088 | 0.831 |
| 3 | 5.0968 | 0.078 | 0.279 | 15442 | 0.791 |
| 4 | 5.4466 | 0.095 | 0.309 | 16502 | 0.764 |
| 5 | 5.7803 | 0.105 | 0.324 | 17513 | 0.772 |
| 6 | 6.8966 | 0.117 | 0.342 | 20894 | 0.874 |
| 7 | 6.9930 | 0.135 | 0.368 | 21187 | 0.823 |
| 8 | 7.4516 | 0.153 | 0.392 | 22576 | 0.824 |

The table shows a clear increase in pressure-head difference as the flow rate rises. This is consistent with the acceleration of water through the restricted orifice area: a larger flow rate requires a higher velocity and therefore produces a greater pressure drop. Although the discharge coefficient does not remain exactly constant, its variation is relatively limited compared with the large change in Reynolds number and pressure head. This indicates reasonably stable meter behavior over the tested turbulent-flow range.

Figure 4: Plot of Pressure Head Difference vs. Volumetric Flow Rate

Figure 4 below shows that the pressure-head difference increases continuously as the volumetric flow rate increases. The relationship is clearly nonlinear and curves upward rather than forming a straight line.For an orifice meter, theory predicts approximately ΔH ∝ Q² because the pressure difference is associated with the change in velocity head. A power fit to the present data gives approximately ΔH = 0.4763Q1.7147 when ΔH is expressed in cm H2O and Q is plotted as Q × 104, with R2 = 0.9573 on the logarithmic power fit with a parabolic curve. The exponent is somewhat lower than the ideal value of 2, so the experimental data do not follow the theoretical square law perfectly; nevertheless, the strong nonlinear increase is consistent with the expected behavior of an orifice meter.

Figure 5: Plot of square root of pressure head difference vs volumetric flow rate

Figure 5 linearizes the pressure-flow relationship by plotting the square root of the pressure-head difference against the volumetric flow rate. The points lie close to a straight line, and a linear fit gives R² = 0.9619. This agrees well with the theoretical orifice meter relation            Q ∝ ΔH. Therefore, this graph is the most useful calibration plot in the experiment: once the calibration line is established, an unknown flow rate can be estimated directly from a measured pressure-head difference without repeating a volumetric collection measurement. The small scatter around the line represents the combined effects of reading uncertainty, flow fluctuations, and other non-ideal experimental conditions.

Figure 6: Plot of discharge coefficient vs pressure head difference

Figure 6 shows the discharge coefficient as a function of pressure-head difference. Although the trendline rises slightly with increasing ΔH, the individual points fluctuate above and below the line and do not show a clear monotonic relationship. For example, Cₒ increases to 0.831 at the second observation, falls through several intermediate observations, reaches 0.874 at the sixth observation, and then returns to about 0.82. Thus, the coefficient can be considered approximately stable with respect to pressure head over the investigated range, but not perfectly constant. The scatter is important because Cₒ is calculated from both measured flow rate and pressure difference; small errors in either quantity is therefore reflected directly in the calculated coefficient.

Figure 7: Plot of discharge coefficient vs square root of pressure head difference

Figure 7 gives a similar conclusion when the discharge coefficient is plotted against √ΔH. The points remain grouped within the same relatively narrow coefficient range, while the fitted line shows only a modest upward tendency. Because √ΔH is directly proportional to flow rate in the ideal orifice equation, the absence of a strong dependence of Cₒ on √ΔH supports the use of an approximately constant coefficient for calibration over the tested turbulent-flow conditions. However, the visible scatter again indicates that using a single coefficient should be regarded as an approximation rather than an exact description of every observation.

Figure 8: Plot of discharge coefficient vs Reynolds number

Figure 8 presents the variation of the discharge coefficient with Reynolds number. All Reynolds numbers are greater than 10,000, confirming that the experiment was performed entirely in the turbulent-flow region. The coefficient remains within 0.751–0.874, with an average of about 0.804. The trendline has a slight positive slope rather than being perfectly horizontal, so the present data should not be interpreted as showing complete independence from Reynolds number. However, there is no strong systematic change across the tested range, and most of the variation appears as scatter around a nearly constant level. This behavior is broadly consistent with the expectation that the discharge coefficient of a sharp-edged orifice becomes only weakly dependent on Reynolds number once the flow is fully turbulent.

Taken together, the five graphs support the basic theoretical behavior of the orifice meter. The pressure-head difference increases nonlinearly with flow rate, whereas √ΔH varies approximately linearly with Q, providing a practical calibration relation. The discharge coefficient is reasonably stable in the fully turbulent regime but exhibits noticeable experimental scatter. Its average value of approximately 0.804 is higher than the commonly quoted range of about 0.60–0.65 for an ideal sharp-edged orifice[4]. Possible causes include slight rounding or wear of the orifice edge, the location of the downstream pressure tapping relative to the vena contracta[4], trapped air or reading uncertainty in the manometer lines, small fluctuations in pump discharge, and timing uncertainty during volumetric flow measurement. These effects can change the measured pressure difference or the actual jet contraction and therefore shift the calculated value of Cₒ. Consequently, the calibration curve obtained from the experimental data is more appropriate for this particular laboratory meter than relying only on a nominal textbook coefficient.

# 7.Conclusion

In this experiment, the performance of a sharp-edged orifice meter was studied and calibrated by relating the experimentally measured volumetric flow rate to the pressure-head difference developed across the orifice. The volumetric flow rate increased from 3.4941 × 10⁻⁴ m³/s to 7.4516 × 10⁻⁴ m³/s, while the corresponding pressure-head difference increased from 0.041 m H₂O to 0.153 m H₂O. This behavior agrees with the expected operating principle of an orifice meter, where an increase in flow velocity through the restricted area produces a larger pressure drop.

The calculated Reynolds number varied from 10586 to 22576 for the eight observations. Since all of these values are well above 4000, the water flow remained fully turbulent throughout the experiment. Under this condition, the coefficient of orifice was expected to remain approximately constant. The experimental coefficient varied between 0.751 and 0.874, with an average value of about 0.804. Although some fluctuation was observed between the individual readings, the relatively narrow range indicates that the meter produced a reasonably consistent response over the investigated flow conditions.

The experimental results therefore confirm the fundamental relationship between flow rate and differential pressure used in orifice-meter measurements. The increase in pressure-head difference with increasing discharge, together with the approximately stable coefficient of orifice in the turbulent region, shows that a practical calibration relationship can be established for the meter. Minor deviations in the coefficient may be attributed to experimental uncertainties such as fluctuations in the pump discharge, difficulty in reading the manometer levels, air bubbles in the pressure lines, stopwatch timing error, or slight wear and rounding of the orifice edge.

Overall, the experiment successfully demonstrated the working principle and calibration of the orifice meter. Once the calibration curve and representative coefficient of orifice are known, the device can be used as a simple and economical method for estimating unknown flow rates from measured pressure differences under similar operating conditions.

# Reference

[1]	R. C. Baker, Flow Measurement Handbook: Industrial Designs, Operating Principles, Performance, and Applications, 2nd ed. Cambridge University Press, 2015. doi: 10.1017/CBO9781107054141.

[2]	W. L. McCabe, J. C. Smith, and P. Harriott, “Unit operations of chemical engineering. Hauptbd.,” 7. ed., in McGraw-Hill’s chemical engineering series. , Boston, Mass.: McGraw-Hill Higher Education, 2005.

[3]	E. Staff, “What is an Orifice Meter? - Instrumentation Tools,” Inst Tools. Accessed: Aug. 21, 2026. [Online]. Available: https://instrumentationtools.com/what-is-an-orifice-meter/

[4]	Y. A. Çengel and J. M. Cimbala, Fluid mechanics: fundamentals and applications, Fourth edition. New York, NY: McGraw-Hill Education, 2017.

[5]	E. J. Finnemore and J. B. Franzini, Fluid Mechanics with Engineering Applications, Tenth edition. in McGraw-Hill’s AccessEngineeringLibrary. New York, N.Y: McGraw Hill LLC, 2002.

[6]	R. L. Mott and J. A. Untener, Applied fluid mechanics, Seventh edition, Global edition. in Always learning. Boston Munich: Pearson, 2016.

[7]	D. W. Green and M. Z. Southard, Perry’s Chemical Engineers’ Handbook, 9th Edition, 9th edition. New York, N.Y: McGraw-Hill Education, 2019.

# Nomenclature

| Symbol | Description | Unit (SI) |
| --- | --- | --- |
| Q | Volumetric flow rate | m³/s |
| Qexp | Experimental volumetric flow rate | m³/s |
| Qtheo | Theoretical volumetric flow rate | m³/s |
| ΔH | Pressure head difference across orifice | m |
| R | Manometer reading difference | m |
| Ao | Area of orifice | m² |
| A | Cross-sectional area of pipe | m² |
| Do | Diameter of orifice | m |
| D | Diameter of pipe | m |
| v | Average velocity of fluid in pipe | m/s |
| Co | Coefficient of orifice | (dimensionless) |
| Re | Reynolds number | (dimensionless) |
| ρ | Density of fluid | kg/m³ |
| ρₘ | Density of manometer fluid (CCl₄) | kg/m³ |
| μ | Dynamic viscosity of fluid | N·s/m² |
| g | Acceleration due to gravity | m/s² |
| V | Volume of collected fluid | m³ |
| t | Time interval for collection | s |
| P1 | Pressure at upstream section | N/m² |
| P2 | Pressure at downstream section | N/m² |
| z₁ | Elevation head at upstream section | m |
| z₂ | Elevation head at downstream section | m |

# Appendix

Sample Calculation for observation 4:

Room temperature = 30℃

Water temperature = 30℃

Density of water at 30℃, ρw= 995.65 kg/m3 [7]

Density of CCl4 at 30℃, ρCCl4= 1574 kg/m3  [7]

Viscosity of water at 30℃, μw= 0.797 mPa.s  [7]

Flow meter volume of water, V = 5 L = 5 × 10-3 m3

Time interval for one cycle of water flow, t = 9.18 s

Volumetric flow rate of water, Qexp = Vt = 5 × 10-3 m39.18 s = 5.4466 × 10−4 m3/s

Internal diameter of the pipe (2-inch, Schedule No. 40), D = 2.067 inch = 0.0525 m [6]

Area of pipe, A = π × D24 = 3.1416 × 0.052524 m2 = 21.65 × 10-4 m2

Average velocity of water in pipe, v = QexpA = 5.4466 × 10-4 m3/s21.65 × 10-4 m2 = 0.2516 m/s

Reynolds number, Re = Dvρwμw = 0.0525 ×0.2516 ×995.650.797 × 10-3 = 16502

∴ The flow regime is turbulent for this observation of water flow.

For 4th observation,

Left reading = 37.2 cm CCl4

Right reading = 20.8 cm CCl4

Manometer reading difference, R = (37.2–20.8) cm CCl4 = 16.4 cm CCl4 = 0.164 m CCl4

Orifice diameter, Do = 1 inch = 0.0254 m

Area of orifice, Ao = π × Do24  = 3.1416 × 0.025424 m2 = 5.067 × 10-4 m2

Pressure Head difference, ∆H=R(ρCCl4-ρw)ρw m H2O

= 0.164×(1574 -995.65)995.65 m H2O

= 0.095 m H2O

Square Root of Pressure Head Difference, ∆H= 0.095 (m H2O)0.5  = 0.309 (m H2O)0.5

Theoretical flow rate, Qtheo = A02g∆H1-DoD4

= 5.067×10-4×2×9.81×0.0951-0.02540.05254  m3/s

= 7.125 × 10−4 m3/s

Coefficient of discharge, Co = QexpQtheo

= 5.4466 × 10-4 m3/s7.125 × 10-4 m3/s

= 0.764

Marking scheme: Long report on ‘Calibration of Orifice Meter’

Name: Mohammad Shahriar Hassan Sifat

Student number: 2302004

| Assessment Component | Weight (%) | Obtained Marks |
| --- | --- | --- |
| Summary | 10% |  |
| Introduction and Theory | 15% |  |
| Experimental Set up | 15% |  |
| Result and Discussion and Conclusion | 25% |  |
| Reference and Nomenclature | 05% |  |
| Quality of Tables | 10% |  |
| Quality of Figures | 10% |  |
| Overall Presentation | 10% |  |
| Total | 100% |  |