Bangladesh University of Engineering and Technology

ChE206: Chemical Engineering Laboratory I

A Long Report on

Experiment No. 4: Calibration of Orifice Meter

Prepared For:

Mahmud Hayet Rudro

Lecturer, Department of Chemical Engineering

Bangladesh University of Engineering and Technology, Dhaka-1000

Prepared By:

Md. Aniruddha

Student ID: 2302038

Section: A2

Lab Group: 2

Partners’ ID:

2302036

2302037

2302039

2302040

Date of Performance: 15 July, 2026

Date of Submission: 29 July, 2026

Statement of Academic Integrity

By signing this experiment, I hereby declare that the work on the experiment titled “Calibration of Orifice Meter” has been prepared honestly and responsibly based on the experimental work, observations, calculations and analysis carried out during the laboratory session. All information, equations, figures, and data obtained from books or other external sources have been properly acknowledged and cited. No data have been fabricated, altered, or intentionally misrepresented, and the interpretations presented in this report reflect the actual experimental findings to the best of my knowledge. I accept full responsibility for the accuracy and originality of the work submitted.

Name: Md. Aniruddha

Student ID: 2302038

Date: August 15, 2026

Signature

Summary

The primary objective of this experiment was to calibrate a sharp-edged concentric orifice meter by evaluating its discharge coefficient across various flow rates. The operation of the device is grounded in the continuity equation and Bernoulli’s principle, where a localized cross-sectional restriction accelerates the flowing fluid, converting static pressure energy into kinetic energy. The resulting differential pressure head (Δ𝐻) was measured across the upstream and downstream pressure taps of a horizontal pipe using a U-tube differential manometer filled with carbon tetrachloride (CCl4), while clean water was circulated as the working fluid by a centrifugal pump.  Because ideal models overlook energy losses and friction, an orifice coefficient (Co) is applied to reconcile theory with actual performance. During testing, experimental volumetric flow rates (QExp) spanned 3.77×10-4 m³/s to 7.54×10-4 m³/s with Reynolds numbers between 11407.88 to 22850.17, confirming fully developed turbulence where actual discharge naturally lagged behind theoretical values due to viscous drag, wall friction, and turbulence. Graphical analysis revealed a strong linear correlation (R²= 0.99947) between √ΔH and QExp alongside relatively flat C0 trendlines that proved stability under turbulent conditions, though the average experimental Co of 0.790 slightly exceeded the standard theoretical range (0.60 to 0.65) due to pump-induced pressure oscillations and slight edge rounding on the orifice face. Ultimately, this successful calibration yielded a dependable empirical reference curve, verifying the meter's reliability and accuracy for industrial flow measurement applications.

Acknowledgement

I sincerely thank the Department of Chemical Engineering at Bangladesh University of Engineering and Technology (BUET) for providing the necessary laboratory facilities and academic support to conduct this experiment under the course ChE 206: Chemical Engineering Laboratory I.

I extend my deepest gratitude to my course instructor, Mahmud Hayet Rudro, Lecturer, for his insightful guidance, clear directions, and continuous support during the experiment and writing process. His advice greatly helped me to analyze the data effectively and successfully.

I also appreciate the support of the laboratory demonstrators in ensuring safe equipment operation and accurate data collection. Lastly, I am thankful for the teamwork and cooperation of my group members during the experimental trial and calculations.

Md. Aniruddha

Student ID: 2302038

Section: A2

Group No: 02

Level-2 Term-2

Department of Chemical Engineering, BUET

Table of Contents

Statement of Academic Integrity	i

Summary	ii

Acknowledgement	iii

Table of Contents	iv

Table of Illustrations	v

1.	Introduction	1

2.	Theory	2

2.1	Orifice Meter & It’s Working Principle	2

2.2	Vena Contracta	3

2.3	Continuity Equation	3

2.4	Bernoulli’s Equation	3

2.5	Differential Manometer Relation	4

2.6	Reynold’s Number	4

2.7	Coefficient of Orifice	5

2.8	Calibration of Orifice Meter	6

3.	Experimental Section	7

3.1	Apparatus	7

3.2	Experimental Setup	7

3.3	Experimental Procedure	8

4.	Result & Discussion	9

4.1	Calibration Results of the Orifice Meter	9

4.2	Calibration Curves of the Orifice Meter	10

4.3	Correlation between experimental and theoretical value	15

4.4	Experimental Errors and Their Consequences	15

5.	Conclusion	16

Reference	17

Nomenclature	18

Appendices	19

Appendix A	19

Observed Data	19

Appendix B	20

Calculated Data	20

Appendix C	21

Sample Calculation	21

Table of Illustrations

List of Tables

Figure 1: Schematic diagram of an orifice meter	2

Figure 2: Schematic diagram of orifice meter apparatus	7

Figure 3: Change of pressure drop DH in the fluid with volumetric flow rate Q	10

Figure 4: Change of square root of pressure drop (√ΔH )with volumetric flow rate (Q)	11

Figure 5: Change in orifice coefficient (Co) with pressure head difference(ΔH)	12

Figure 6: Change in orifice coefficient (Co)with square root of pressure head (√ΔH)	13

Figure 7: Change in orifice coefficient (Co)with Reynolds Number	14

List of Tables

Table 1: Observed data for volume of water, time intervals and manometer readings	9

Table 2: Observed Data for Volume of Water, Time Interval and Manometer Reading	19

Table 3: Calculated data for Volumetric Flow Rate, Pressure Head Difference, Velocity, Reynolds number and Orifice Coefficient for different observations	20

# Introduction

Measuring fluid flow is essential in industries like water supply, oil refining, chemical processing, and power generation to oversee operations, guarantee safety, control quality, and track material balances. This is typically done by installing dedicated instruments inside pipelines to calculate the mass or volume of fluid moving through a cross-section over time. Differential pressure meters, specifically orifice meters, are especially popular for this purpose because they are inexpensive, robust, lack moving parts, and feature a straightforward design.

An orifice meter functions by placing a flat plate containing a centered circular hole perpendicular to the flowing fluid. This restriction narrows the flow path, causing the fluid to speed up. Based on Bernoulli's principle and the continuity equation, this acceleration leads to a static pressure drop just past the plate. The fluid streamlines keep converging beyond the plate, achieving their narrowest point and highest speed at the vena contracta. The resulting pressure difference between the upstream pipe and the restricted section is then recorded using a differential U-tube manometer containing a dense fluid like carbon tetrachloride.

Because real-world fluids encounter wall turbulence, friction, viscous drag, and flow separation, theoretical flow calculations based strictly on Bernoulli's principle tend to overestimate actual volume. To account for these energy losses, an empirical value called the orifice coefficient (Co) is applied to the theoretical equations.

As long as fluid density and viscosity are factored into the Reynolds number, flow rates can be determined directly from pressure head measurements without manual collection. Even so, practical challenges often arise, such as edge rounding on the orifice or pressure fluctuations from pumps, which can alter the discharge coefficient and affect measurement accuracy.

# Theory

## Orifice Meter & It’s Working Principle

An orifice meter is a differential pressure type flow measuring device used to determine the volumetric flow rate of a fluid flowing through a pipeline. It consists of a thin plate with a sharp-edged circular opening, called an orifice, which is installed perpendicular to the direction of flow. A schematic diagram of a typical orifice meter is shown in Figure 1[1]

Figure 1: Schematic diagram of an orifice meter

When fluid flows through the pipe and passes through the orifice, the effective flow area suddenly reduces. Due to this reduction in area, the velocity of the fluid increases. According to Bernoulli’s principle, an increase in velocity is accompanied by a decrease in static pressure. As a result, a pressure difference is created between the upstream section and the downstream section of the orifice. This pressure difference is measured using a differential manometer and is directly related to the flow rate of the fluid. Thus, by measuring the pressure drop across the orifice, the volumetric flow rate can be determined.

The working principle of the orifice meter is therefore based on the conversion of pressure energy into kinetic energy, as described by the continuity equation and Bernoulli’s equation.[2]

## Vena Contracta

When fluid approaches an orifice plate, the streamlines converge toward the opening. Due to fluid inertia, the fluid stream cannot execute abrupt 90-degree turns around the sharp edges of the restriction, causing the converging streamlines to continue contracting even after exiting the orifice. The specific location downstream where this fluid jet reaches its absolute minimum cross-sectional area and maximum velocity is defined as the vena contracta. Downstream of the vena contracta, the fluid expands to refill the pipe diameter, creating severe flow separation, recirculation eddies, and turbulence that dissipate mechanical energy into unrecoverable heat, resulting in permanent head loss. Because of these irreversible fluid friction losses, practical flow calculations require the incorporation of an empirical discharge coefficient.[3]

## Continuity Equation

For steady, incompressible flow, the continuity equation is expressed as,

A1v1 = A2v2,

Where,

A=cross-sectional area

v=denotes the average velocity.

In an orifice meter, the flow path narrows at the restriction, which, according to the continuity equation, causes the fluid velocity to rise at that location. This acceleration results in a drop in static pressure. Consequently, the continuity equation accounts for the velocity increase responsible for generating the pressure differential utilized to measure flow.

## Bernoulli’s Equation

For steady, incompressible, and frictionless flow between two sections along a streamline, Bernoulli’s equation is expressed as

P1ρg+v122g+z1=P2ρg+v222g+z2

For a horizontal pipe, z1=z2

Hence,

P1-P2ρg=v12-v222g

Here, P1 and P2 represent the static pressures at sections 1 and 2, respectively. For a horizontal pipe, the elevation terms (z1 & z2) cancel out, and the pressure difference is balanced by the change in velocity head. This relation forms the theoretical basis for pressure-based flow measurement devices such as the orifice meter.

## Differential Manometer Relation

The pressure difference across the orifice is measured using a U-tube manometer containing carbon tetrachloride (CCl₄). The equivalent pressure head in terms of water is given by

ΔH=R(ρCCl4ρw-1)

where,

R = manometer reading difference

𝜌𝐶𝐶𝑙4 = density of manometer fluid

𝜌𝑤 = density of water

In an orifice meter, the sudden reduction in flow area causes an increase in velocity and a corresponding drop in static pressure across the orifice. This pressure difference between the upstream and downstream sections is measured using a differential manometer. The manometer reading is then converted into an equivalent pressure head in terms of water, which is directly used to determine the discharge through the orifice.

## Reynold’s Number

The Reynolds Number(Re) is a dimensionless parameter in fluid mechanics representing the ratio of inertial forces to viscous forces within a moving fluid.

Re=ρvDμ

where,

𝜌 = density of fluid

𝑣 = average velocity

𝐷 = pipe diameter

𝜇 = dynamic viscosity of fluid 5

In the context of orifice meter calibration, the Reynolds number determines the stability of the discharge coefficient (𝐶o). In laminar or transitional flow regimes, 𝐶o varies significantly with changes in velocity because boundary layer thickness changes with flow rate. However, in fully turbulent flow (Re>4,000), inertia forces dominate over viscous forces, stabilizing jet contraction at the vena contracta. As a result, 𝐶o becomes nearly constant and independent of Reynolds number variations.

## Coefficient of Orifice

The derived theoretical equation of discharge assumes a perfect, frictionless fluid flow. However, real-world fluid dynamics says that a perfect fluid flow is not possible, some amount of energy loss will always take place. In real engineering systems, the measured flow rate (QExperimental) will always fall short of the theoretical value. This discrepancy happens for frictional loss and fluid flow contraction.

In order to convert theoretical flow rate to real flow rate, the Coefficient of Orifice (Co) is introduced. This factor is the ratio of the actual flow measured empirically to the theoretical flow rate:

Co=QExperimentalQTheoretical

Substituting QTheoretical,

Co=QExperimentalAo2gΔH1-DoD4

In practice, QActual is determined physically during the experiment by using a flowmeter. If theoretical discharge is calculated using manometer readings (ΔH) and the equation derived earlier, Co can be directly determined.

## Calibration of Orifice Meter

Calibration connects theoretical fluid mechanics with the real-world performance of an orifice meter. Because factors like fluid velocity, fluid properties, and unique pipe geometry introduce real-world losses that theory cannot predict, an uncalibrated meter cannot deliver accurate measurements. During calibration, fluid is pumped through the system at multiple steady flow rates. For each rate, the actual discharge (measured via a reference flowmeter) and the corresponding pressure head difference across the orifice plate (recorded via a manometer) are gathered simultaneously. Using this data, the discharge coefficient is calculated for each flow rate and plotted against Reynolds numbers, pressure head differences, and the square root of those differences to generate calibration curves. Once completed, these curves allow the orifice meter to be used reliably for measuring unknown flow rates in practical engineering applications.

# Experimental Section

## Apparatus

- Orifice plate

- Differential manometer

- Centrifugal pump

- Reservoir

- Flowmeter

- Stopwatch

- Thermometer

- Scale

## Experimental Setup

Figure 2: Schematic diagram of orifice meter apparatus

The experimental setup consists of a concentric orifice meter for flow measurement, a centrifugal pump for fluid circulation, a differential U-tube manometer for recording pressure head differences, and a flow meter to measure discharge. Water is pumped from a reservoir through a pipeline containing the orifice plate, where an inlet bypass line is installed to regulate the flow rate and prevent overflow. The entrance valve is used to establish the maximum flow limit, while the outlet control valve is adjusted to stabilize and set specific flow rates. Pressure tapping points located upstream and downstream of the orifice plate are 8 connected directly to the differential U-tube manometer filled with carbon tetrachloride (CCl4) as the indicating liquid.

## Experimental Procedure

Initially, the supply valve was opened to fill the storage reservoir with water to provide a continuous fluid feed to the centrifugal pump. The entrance valve and the bypass valve were then adjusted to establish a steady, uniform flow rate while defining the upper flow limit and preventing tank overflow. Prior to recording measurements, the differential U-tube manometer was inspected to ensure that both fluid limbs were balanced at the same initial datum level. Once secure pressure tap connections were confirmed, the control valve was adjusted to set the initial flow rate, allowing sufficient time for the flow to stabilize before measurements were taken. To determine the experimental discharge, a fixed volume of water corresponding to half a cycle of the flowmeter (5 L) was collected, and the time interval was measured using a stopwatch. Under identical operating conditions, the liquid column heights in the left and right limbs of the manometer were recorded to evaluate the differential pressure head across the orifice plate. This measurement sequence was repeated for seven additional runs at incrementally higher flow rates. Finally, both the room temperature and water temperature were measured using a thermometer to determine the physical properties (density and viscosity) of the carbon tetrachloride and water.

# Result & Discussion

## Calibration Results of the Orifice Meter

| Observation | Volumetric flow rate Q×104 (m3/s) | Velocity V (m/s) | Reynolds number (Re) | Pressure Drop, ΔH (m H2O) | Coefficient of orifice meter, C0 |
| --- | --- | --- | --- | --- | --- |
| 1 | 3.77 | 0.174 | 11407.88 | 0.041 | 0.809 |
| 2 | 4.36 | 0.201 | 13208.08 | 0.057 | 0.789 |
| 3 | 4.76 | 0.220 | 14428.25 | 0.071 | 0.777 |
| 4 | 5.35 | 0.247 | 16220.2 | 0.090 | 0.772 |
| 5 | 5.97 | 0.276 | 18099.96 | 0.111 | 0.779 |
| 6 | 6.87 | 0.318 | 20809.98 | 0.134 | 0.812 |
| 7 | 7.14 | 0.330 | 21642.38 | 0.153 | 0.791 |
| 8 | 7.54 | 0.349 | 22850.17 | 0.172 | 0.788 |

Table 1: Observed data for volume of water, time intervals and manometer readings

The data presented in Table 1 summarizes the experimental results obtained during the calibration of the orifice meter across eight distinct flow rates. The table documents the progression of the actual volumetric flow rate (Q), which was systematically increased from 3.77×10-4 m³/s to 7.54×10-4 m³/s. As the flow rate increased, the measured pressure head difference (ΔH) exhibited a corresponding rise from 0.041 m H₂O to 0.172 m H₂O, consistent with the principles of fluid dynamics outlined in the theoretical model.

The calculated Reynolds numbers (Re) for the system span from 11407.88 to 22850.17, confirming that the fluid flow remained entirely within the turbulent regime during all eight observations. Finally, the empirically derived Coefficient of Orifice (Co) is recorded for each trial, showing values that fluctuate between 0.809 and 0.788. This tabulated dataset provides the foundational values necessary for generating subsequent calibration curves and evaluating the performance characteristics of the meter.

## Calibration Curves of the Orifice Meter

Figure 3: Change of pressure drop DH in the fluid with volumetric flow rate Q

The graph illustrates the relationship between the differential pressure head (Δ𝐻) and the volumetric flow rate (𝑄). It can be observed that an increase in the volumetric flow rate results in a corresponding increase in the pressure difference across the orifice plate. The experimental data points form an upward-curving parabolic trend, indicating a nonlinear relationship between the two variables. This behavior agrees with the theoretical operating principles of an orifice meter, where the pressure differential is proportional to the square of the volumetric flow rate, expressed as Δ𝐻 ∝ 𝑄². Thus, the graphical results are consistent with the expected theoretical behavior of flow through an orifice restriction.

Figure 4: Change of square root of pressure drop (√ΔH )with volumetric flow rate (Q)

Figure 3 illustrates the relationship between the square root of the differential pressure head (√Δ𝐻) and the volumetric flow rate (𝑄). According to the theoretical principles governing the operation of an orifice meter, the volumetric flow rate is directly proportional to the square root of the pressure drop across the orifice plate, expressed as 𝑄 ∝ √Δ𝐻. Therefore, when 𝑄 is plotted against √Δ𝐻, the experimental data are expected to exhibit an approximately linear relationship, resulting in a straight-line trend.

Figure 5: Change in orifice coefficient (Co) with pressure head difference(ΔH)

The graph illustrates the relationship between the orifice coefficient (𝐶ₒ) and the differential pressure head (Δ𝐻). According to the theoretical behavior of an orifice meter, for sufficiently high Reynolds numbers, the discharge coefficient of an orifice meter becomes relatively insensitive to further increases in Reynolds number. Since all experimental observations were within the turbulent regime, 𝐶ₒ is expected to remain approximately constant. Consequently, in a region of fully developed turbulent flow, the coefficient is expected to remain relatively stable. The experimental trendline becomes progressively more horizontal with increasing pressure head difference, which supports this theoretical behavior. Although the overall variation in 𝐶ₒ is relatively small, two noticeable outliers are present, with values of 0.809 and 0.812. Despite these deviations, the orifice coefficient generally remains close to 0.8 throughout the experiment.

Figure 6: Change in orifice coefficient (Co)with square root of pressure head (√ΔH)

The graph illustrates the relationship between the orifice coefficient (𝐶ₒ) and the square root of the differential pressure head (√Δ𝐻). The plotted data exhibit some degree of scatter, with the measured 𝐶ₒ values ranging approximately from 0.772 to 0.812. Despite this variation, the fitted linear trendline is nearly horizontal, showing only a slight negative slope. This behavior agrees with the theoretical expectation that the orifice coefficient remains approximately constant under turbulent flow conditions.

Figure 7: Change in orifice coefficient (Co)with Reynolds Number

The graph illustrates the relationship between the orifice coefficient (𝐶ₒ) and the Reynolds number (𝑅ₑ). The experimental values of 𝐶ₒ show some scatter, ranging approximately from 0.772 to 0.812 over the Reynolds number range 11408 to 22850. However, most of the values remain close to 0.80, indicating that the discharge coefficient does not undergo significant variation within the tested flow conditions. As the Reynolds number increases, the trendline becomes relatively flatter, suggesting that 𝐶ₒ approaches a nearly constant value at higher Reynolds numbers. This behavior is consistent with the theoretical characteristics of orifice meters, where the discharge coefficient becomes relatively insensitive to Reynolds number once the flow is sufficiently turbulent

Several experimental factors may have contributed to the deviations from the expected theoretical behavior. These effects are particularly significant at lower flow rates, where small measurement uncertainties can produce relatively larger errors. Fluctuations in the manometer fluid level may have affected the accuracy of the pressure-head measurements, while variations in pump discharge could have caused instability in the flow rate. In addition, minor leakage from the system before the outlet may have resulted in a slight underestimation of the actual discharge. To minimize these errors in future experiments, the pump should be operated at a stable delivery pressure, all trapped air should be completely removed from the piping and measurement lines before taking readings, and more precise timing and differential-pressure measuring instruments should be employed.

## Correlation between experimental and theoretical value

The experimental discharge was consistently lower than the theoretical discharge. This reduction is represented by the discharge coefficient being less than unity. The deviation between ideal theoretical predictions and actual experimental can be caused by two main fluid phenomena:

1. Fluid Contraction: The cross-sectional area of the fluid jet contracts to a minimum smaller area than the physical orifice opening downstream of the plate.

2. Energy Dissipation: Mechanical energy is converted into heat due to fluid viscosity, eddy formation, and sudden flow expansion downstream.

While standard theoretical values for sharp-edged orifice meters typically range from 0.60 to 0.65, the experimental values obtained in this test averaged 0.790. This shift reflects specific mechanical, such as blunt or rounded edges of the orifice plates, characteristics of the experimental apparatus as the orifice meter is using for long time.

## Experimental Errors and Their Consequences

During the experiment, several sources of experimental error were identified, which influenced the accuracy and consistency of the obtained results. The most significant of error was associated with the differential manometer. As the accuracy of the scale is ±0.5cm, we can’t measure more accurate than it.

Another important source of uncertainty can be caused by the shape and geometry of the orifice plate. Slight rounding or wear on the upstream edge of the orifice plate reduces jet contraction. A reduced contraction increases actual flow area at the vena contracta, causing the measured 𝐶o values to be higher than standard sharp-edged theoretical values.

# Conclusion

The orifice meter calibration was successfully carried out to find its discharge coefficient empirically and define the core relationship connecting pressure head variation to actual volumetric flow. Through the use of Bernoulli's equation and continuity principles, standard theoretical discharge models were weighed against physical flow rates gathered across eight unique operating points.

These findings successfully confirmed the underlying fluid dynamics principles of the device. The results demonstrated that actual flow increases parabolically relative to the pressure head difference, a pattern that became linear when mapped against the square root of the pressure drop. Throughout the testing, flow remained completely turbulent, with Reynolds numbers spanning from 11407.88 to 22850.17. Under these conditions, the computed orifice coefficients fell between 0.809 and 0.788.

Even though theory dictates that this coefficient should stay completely flat and constant during high turbulence, the calibration plots showed visible scatter and strayed from this ideal trend. Analysis revealed that these discrepancies stemmed from real-world experimental constraints, such as pump-induced vibrations disturbing the manometer liquid, human stopwatch delays, and natural energy losses from friction and turbulent eddies around the plate.

In spite of these operational energy losses, the resulting calibration curves effectively captured the unique performance profile of the equipment. This empirical data offers a trustworthy benchmark for the meter's behavior, successfully meeting the experiment's main goal and allowing the device to be used reliably for finding unknown flow rates in future practical work.

Reference

[1] Instrumentation Tools. (2023). What is an orifice meter? Instrumentation Tools. Retrieved from https://instrumentationtools.com/what-is-an-orifice-meter/

[2] W. L. McCabe, J. C. Smith, and P. Harriott, Unit Operations of Chemical Engineering. McGraw-Hill, 2005.

[3] Cengel, Y. A., & Cimbala, J. M. (2017). Fluid mechanics: Fundamentals and applications (4th ed.). McGraw-Hill Education.

[4] W. Wagner and A. Pruss, “The IAPWS Formulation 1995 for the Thermodynamic Properties of Ordinary Water Substance for General and Scientific Use,” J. Phys. Chem. Ref. Data, vol. 31, no. 2, pp. 387–535, 2002, doi: 10.1063/1.1461829.

[5] D. W. Green and R. H. Perry, Perry’s Chemical Engineers’ Handbook, Eighth Edition. McGraw Hill Professional, 2007.

Nomenclature

| Symbol | Meaning | Unit (SI) |
| --- | --- | --- |
| A | Cross-sectional area of the pipe | m² |
| Ao | Cross-sectional area of the orifice | m² |
| Co | Coefficient of Orifice | Dimensionless |
| D | Internal diameter of the main pipe | m |
| Do | Orifice diameter | m |
| g | Acceleration due to gravity | m/s² |
| PA, PB | Static pressure at points A and B | Pa |
| QExperimental | Experimental volumetric flow rate | m³/s |
| QTheoretical | Theoretical volumetric flow rate | m³/s |
| R | Manometer reading difference | m |
| Re | Reynolds Number | Dimensionless |
| t | Time interval | s |
| V | Volume of water | m³ |
| v | Average fluid velocity | m/s |
| vA, vB | Fluid velocities at points A (upstream) and B (vena contracta) | m/s |
| zA, zB | Elevation heads | m |
| ΔH | Pressure head difference | m |
| μw | Dynamic viscosity of water | Pa⋅s |
| ρ | Density | kg/m³ |
| ρw | Density of water | kg/m³ |
| ρCCl4 | Density of Carbon Tetrachloride (manometric fluid) | kg/m³ |

Appendices

## Appendix A

### Observed Data

Room Temperature = 30℃

Water Temperature = 31℃

Schedule Number = 40

Pipe Diameter, D = 2.067 inch = 0.0525 m [4]

Orifice Diameter, Do = 1 inch = 0.0254 m

Table 2: Observed Data for Volume of Water, Time Interval and Manometer Reading

| No of Observation | Time Interval, t (s) | Volume of water, V (L) | Manometer Reading |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Left (cm) | Right (cm) | Difference (cm) |
| 1 | 13.28 | 5 | 32.3 | 25.5 | 6.8 |
| 2 | 11.47 |  | 33.8 | 24.2 | 9.6 |
| 3 | 10.50 |  | 35.0 | 23.2 | 11.8 |
| 4 | 9.34 |  | 36.5 | 21.4 | 15.1 |
| 5 | 8.37 |  | 38.3 | 19.8 | 18.5 |
| 6 | 7.28 |  | 40.3 | 17.8 | 22.5 |
| 7 | 7.00 |  | 41.8 | 16.2 | 25.6 |
| 8 | 6.63 |  | 43.5 | 14.7 | 28.8 |

## Appendix B

### Calculated Data

Density of water at 31° C, ρw  = 995.3 kg/m3[4]

Density of CCl4 at 30° C, ρCCl4  = 1590 kg/m3[5]

Viscosity of water at 31° C, µw  = 7.97 ×10-4 Pa.s [4]

Area of Orifice, Ao = 5.067×10-4 m2

Area of Pipe, A = 2.165×10-3 m2

Table 3: Calculated data for Volumetric Flow Rate, Pressure Head Difference, Velocity, Reynolds number and Orifice Coefficient for different observations

| Observation | Volumetric flow rate Q×104 (m3/s) | Velocity V (m/s) | Reynolds number (Re) | Pressure Drop, ΔH (m H2O) | Coefficient of orifice meter, C0 |
| --- | --- | --- | --- | --- | --- |
| 1 | 3.77 | 0.174 | 11407.88 | 0.041 | 0.809 |
| 2 | 4.36 | 0.201 | 13208.08 | 0.057 | 0.789 |
| 3 | 4.76 | 0.220 | 14428.25 | 0.071 | 0.777 |
| 4 | 5.35 | 0.247 | 16220.2 | 0.090 | 0.772 |
| 5 | 5.97 | 0.276 | 18099.96 | 0.111 | 0.779 |
| 6 | 6.87 | 0.318 | 20809.98 | 0.134 | 0.812 |
| 7 | 7.14 | 0.330 | 21642.38 | 0.153 | 0.791 |
| 8 | 7.54 | 0.349 | 22850.17 | 0.172 | 0.788 |

## Appendix C

### Sample Calculation

For observation 3:

Volumetric flow rate:

Volume of water, V = 5.0 × 10-3 m3

Time interval, t = 10.50 s

Volumetric flow rate of water, Q = Vt = 5.0 × 10-3 m310.50 s = 4.76× 10-4 m3/s

Average Velocity:

Pipe Diameter (Schedule No. 40), D = 2.067 inch = 5.25 × 10-2  m

Area of Pipe, A = π D24= π (5.25 × 10-2) 24 m2  = 2.165 × 10-3 m2

Average Velocity of water in pipe, v =  QA  =  4.76 × 10-4 2.165 × 10-3 m/s = 0.220 m/s

Pressure Drop (ΔH in m H2O):

Left reading = 35.0 cm CCL4

Right reading = 23.2 cm CCL4

Manometer reading difference, R = (32.3 – 25.5) cm CCL4 = 11.8 cm CCL4=11.8×10-2 m CCL4

Pressure Drop, ΔH = R × ( ρCCL4ρw -1 )

= 11.8×10-2 × ( 1590995.3  - 1) m H2O = 0.071 m H2O

Co-efficient of Discharge

Orifice Diameter, D0 = 1 inch = 2.54 × 10-2 m

Area of Orifice, Ao = π D024= π (2.54 × 10-2) 24 m2  = 5.067 × 10-4 m2

Coefficient of discharge, Co  = QAo × 2gΔH1-(D0D)4 = 4.76 × 10-4 5.067 × 10-4 × 2×9.8× 0.071 1-(2.54 × 10-2 5.25 × 10-2)4 = 0.777

Reynolds number

Reynolds Number, Re =  ρwvD µw = 995.3  × 0.220 × 0.0525 7.97 ×10-4 = 14428.25

Marking scheme: Long report on ‘Calibration of Orifice Meter’

Name: Md. Aniruddha

Student number: 2302038

| Assessment Component | Weight (%) | Obtained Marks |
| --- | --- | --- |
| Summary | 10% |  |
| Introduction and Theory | 15% |  |
| Experimental Set up | 15% |  |
| Result and Discussion and Conclusion | 25% |  |
| Reference and Nomenclature | 05% |  |
| Quality of Tables | 10% |  |
| Quality of Figures | 10% |  |
| Overall Presentation | 10% |  |
| Total | 100% |  |